AFM Operating Procedure

- Turn on both screens (computer is normally left running with screens turned off)

- Open XEP and XEC from start-up screen:

  - XEC:

    - Turn on the fiber optic light source; set the light control to ~halfway position

    - Click on the “XEC” icon on the desktop

  - XEP:

    - Starting XEP will automatically raise the Z-scanner head to 29000um (max height) and return the sample stage to the home position (0, 0)

    - Load the desired sample onto the stage (I usually use the “TGT-1” or other standard sample to check the tip condition before running “real” samples).

      - In the “XY Stage” window, click on the “Step & Scan” button

      - Make sure that the Step & Scan values are set to 0 and 75000

      - Click the “GoTo” button to move the stage toward the front

      - When stage has stopped, load the sample of interest

      - Click on the “Basic” tab, then click “R. Origin” to return the stage to the home (0, 0) position

    - Open new session: Tools\Session Manager\Create Session

    - Setup\Input\Scan Config

    - Topography is default selection; select Error and Amplitude also

    - Change “Step” interval to 50um (personal preference)

    - Select “Repeat” in the Scan Control window; set to > 20

    - De-select Automatic Image Storing

    - Adjust the position of the red dot on the PSPD by using the small screws located on the scanner head. “A-B” and “C-D” should both be -0.3 < x < +0.3. “A+B” should be > ~2.5V

    - Bring Z-scanner toward sample by clicking and holding in the lower part of the grey region in the “Z-Stage” window. Downward speed increases the further the cursor is from the center line

    - WARNING: Bringing scanner down too far and/or too fast can result in contact with sample surface. This can damage the tip (~$50) and/or the scanner (~$20k). Be Careful!

    - Lower Z-scanner until it is ~3000um above the sample surface. This value will vary depending on the sample. Check with a trained used if not sure what this value is.

    - Use the “Focus Stage” to bring the microscope into focus on the sample surface. XEP allows you to do this without moving the Z-scanner itself to avoid unwanted contact (i.e. crashing into) the sample surface.

    - Make sure the sample surface is centered on the screen. This will ensure that the tip will scan the region of interest.

    - Raise the Focus Stage so that the microscope is focused on the tip. Then lower the Focus Stage ~100um

  - Warm up the scanner for 15 min or more to reduce thermal drift when imaging the “real” sample

    - Set the scan size to the appropriate value (10um for XY LV, 90um for XY HV)

    - Click on the “Image/Start” button to begin the scan. The topography signal should be ~0, since the tip is not in proximity with the sample surface

    - Let scanner run for the desired time, then click “Cancel” to stop the scan

  - Now select the “Approach” button in the Z Stage window to automatically lower the tip to close proximity with the sample. The tip and sample should appear as in the picture below:

    - Select the “Slow Scan Reverse” to ensure that the scan motion is from top to bottom. This reduces the drift in the scan.

    - Click on the “Start” button to begin the scan

    - Adjust the scan parameters (Scan Rate, Z Servo Gain, Set Point) to optimize the image using the trace-retrace topography signals. These should appear identical but may be offset.

    - If not symmetric, adjust the scan parameters so that the trace and retrace topography scans appear the same.

    - If the standard sample shows that the tip is damaged, contact Engineering for tip change

  - When the scan has completed, the tip will continue scanning in the X-direction but not in the Y-direction. You can also stop the scan before it completes the set size by clicking on the “Cancel” button

  - If done scanning, raise the Z-scanner by clicking on the “Lift Z” button. This will drive the scanner to the maximum Z height of 29 000um.

  - Drive the stage toward the front and remove sample

    - In the XYStage window, click on the “Step & Scan” button

    - Make sure that the Step & Scan values are 0 and 75000

    - Click on the GoTo button to move stage

    - Remove the sample

    - Load the next sample, if desired

    - You can leave the stage in this position if desired. This will make it easier to load a sample next time.

  - Close XEP

    - Turn off the laser head by clicking on the Head On/Off icon; answer Yes at the prompt “Do you really want to turn off Head?”

    - Click on the “X” in the upper right corner of the window

  - Close XEC

    - Turn down the lamp intensity

    - Turn off the power switch

    - Close XEC by clicking on the “X” in the upper right corner of the window

  - Close XEP

    - Turn