Procedure for KDF Power Off/Glitch Recovery

- Power Off

  - Check the EPO buttons are released

- Four EPO buttons on the KDF tool

  - RF Power

  - L/L Door

  - Front Main Cover

  - On Auxiliary Cabinet

  - Circuit Breakers

- Auxiliary Power Cabinet

- Rough Pump

- Cryo Compressor

- Check with Facilities if no power

  - Initialize Computer Login

- MANAGER login

  - Cryo Temperature check.

    - Cryo Temp Reading on L/L and on rear of chamber

    - >25K the Cryo must be Regened

  - Check Chamber and L/L Pressures

    - Leak Rates run after recovery

  - Initialize

- Glitch

  - Computer Login

- MANAGER login

  - Cryo Temperature check.

    - Cryo Temp Reading on L/L and on rear of chamber

    - >25K the Cryo must be Regened

  - Check Chamber and L/L Pressures

    - Leak Rates run after recovery

  - Initialize

- Computer Does Not Recover

  - Reset ON/OFF on the lower cover computer (door with key)

Dirk/Ginny   11-Oct-06