Agilent 5900 ICP-OES Protocol

Version 1         9/29/2020

- Front left of instrument (soft power) should be illuminated

  - If not, main power switch on back left needs to be turned on and left to purge optics (min 1 hour?)

- Ensure SPS 4 Autosampler is on and illuminating all green lights in top left

- Ensure Argon is connected and flowing to instrument

- Turn on chiller (18°C)

- Clamp pump tubing down

  - Front white/black tubing = rinse solution

  - Middle white/black tubing = internal std (Y 5ppm)

  - Back blue tubing = waste line

- Open ICP Expert Software

  - Connect instrument (top left) to IP 192.168.1.128

  - Ensure all status lights are green

    - If Camera light is yellow, most likely that means Peltier temp is not going down on its own. Click the “Service Only” button and enter password: Manticore. Go into “Detector” tab and click “Peltier Control On”. Peltier temp should then reduce to –40C. Make sure to log out of service options so nothing else is changed.

- Select a “New From” Worksheet template or create a new one

  - Configuration tab

    - Enable: Autosampler & SVS 6/7 if using autosampler and auto-switching valve

    - Enable QC if using QCs

    - Enable Intelliquant

  - Elements tab

    - Select/verify all element wavelengths of interest

      - Take note of possible interferences in red

    - Select Y and mark as internal std instead of analyte

  - Conditions Tab

    - Under Measurement Conditions --> Viewing Mode --> Select SVDV  for simultaneous radial/axial measurements

    - Sample loop is set at 1mL. If changing, use calculator and change sample loop & it will adjust all other common conditions accordingly

  - Standards Tab

    - Ensure all standard levels are added and include appropriate concentrations for each element

    - Ensure Max Calibration Error is above 10% to be able to see error data

  - QC Tab

    - Ensure QC is selected and approproate concentrations are entered on right side under Limits --> Defined Concentrations

  - Sequence Tab

    - Enter # of samples on right side (can copy & paste from an excel sheet)

    - To enter a QC, right click on sample/cal above it and choose Insert solution below --> CCV

- When analysis is completed, turn off plasma

  - Pull rinse & internal std tubes out of containers & run pump fast to dry out lines

- Turn off chiller

- Unlock pump tubing from peri pump (leave relaxed)

- Cap all samples & standards

- Save data